## Slide 1
In-house Training – 1:
Analysis of Pesticide Residues
QEL, Spices Board
Lecture - 1
Introduction to instrumentation:
LC-MS/MS
03-Jul-23
Dr. Ramesh BN
Spices Board

## Slide 2
Mobile Phase reservoir
Stationery phase
Detection System
A + B
Mixture
Chromatography
HPLC, GC…

## Slide 3
HPLC Analysis
Limitations
When conventional detectors are used, identification of unknowns by retention time matching
Possibility of errors
To avoid these errors, we use mass spectrometry along with chromatography
1
2
3
Inject Standard:
Inject Sample:
R.T. Match
Compound 3 present in sample
Problems:
⤷ Two compounds with same Retention Time: FALSE POSITIVE
⤷ Retention Time shift for one compound: FALSE NEGATIVE

## Slide 4
Mass Spectrometry
Why?
A mass spectrometer (MS) looks at the masses of the analytes being tested
  ⤷ Uses information derived from these masses for identification
  ⤷ High degree of specificity
  ⤷ Retention time can be used for additional confirmation

## Slide 5
MS as a detector
Hyphenated techniques
An MS can function as a detector to a chromatographic system, replacing conventional detectors
  ⤷ Gas chromatography (GC-MS), liquid chromatography (LC-MS)
  ⤷ Can provide very specific and sensitive analysis
  ⤷ Currently, MS is the default detector in trace analysis work
  ⤷ ‘Tandem’ MS is possible, which gives even more specificity and sensitivity (MS/MS)

## Slide 6
Mass Spectrometry
§ General principles

## Slide 7
Mass Spectrometry
Principles - 1
A mass spectrometer can ‘see’ only charged entities (ions)
Typically functions under high vacuum
  ⤷ The ions need a stable path through the mass spectrometer, without colliding with air molecules or with one another
The ‘mean free path’: average distance between collisions for a gaseous ion – must be high
To detector
760 Torr
To detector
10-5 Torr

## Slide 8
Mass Spectrometry
Principles - 2
In MS the vacuum is achieved in two stages:
  ⤷ The rouging pump (outside the instrument): from 760 Torr to 10-3 Torr
  ⤷ The turbo molecular pump (inside the instrument): from 10-3 Torr to 10-5 Torr

## Slide 9
Mass Spectrometry
Principles - 3
The connection between LC and MS is called the ‘Interface’:
  ⤷ From HPLC the mobile phase flows out
  ⤷ MS cannot handle this – it is a vacuum technique
  ⤷ The liquid mobile phase has to be evaporated and taken into the MS
  ⤷ Loss of analytes should be kept to a minimum
  ⤷ In LC-MS, ionization also happens in the interface

## Slide 10
Mass Spectrometry
Principles - 4
MS can operate in two ways.
First way:
  ⤷ Molecule of interest is ionized
  ⤷ Typically called ‘parent’ ion.
  ⤷ Number of parent ions can be counted for quantification (LC-MS).
  Second way:
  ⤷ Molecule of interest is ionized, to get the ‘parent’ ion
  ⤷ Parent ion is broken down into multiple ‘daughter’ ions
  ⤷ Selected ‘daughter’ ions are counted for quantification (LC-MS/MS)
Why is this useful?

## Slide 11
Mass Spectrometry
§ Components and operation

## Slide 12
High vacuum
Ionzation
Select
Parent ion
Make daughter ions
Select daughter ion
Count ions
LC MS/MS
LC MS
In LC MS/MS, ionization happens in atmospheric pressure
Ionization is soft: MH+ ion is formed
Standard spectrum libraries not possible
HPLC
Mass Spectrometry
Components - overview

## Slide 13
Mass Spectrometry
Ionization
From HPLC
pump
Nebulizer Gas
+
+
+
+
+
+
+
+
+
+
+
+
+
Evaporation
Coulomb Explosion
From solvent
Parent ion
Used for MS/MS
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
+
Charged Droplets
N2
vacuum
-ve
High voltage
~ 5500V
Heating gas

## Slide 14
Mass Spectrometry
Quadrupoles (1)
Four electrodes arranged in a hyperboloid fashion
Two Voltages:
  Radio Frequency - RF (alternating)
  Direct Current - DC
  Applied simultaneously, the two voltages will trap ions within the quadrupole area
  Scanning the voltages will filter the ions out by mass:
    For one RF:DC value, only one mass can have a stable trajectory through the quadrupole
    When RF:DC is progressively increased (scanned) Ions pass through in the order of masses
RF

## Slide 15
Mass Spectrometry
Quadrupoles (2)
+
+
-
-
-
-
+
+
Opposite electrodes of the Quadrupole have the same potential
The potential alternates with the same frequency as the RF Voltage

## Slide 16
Mass Spectrometry
Quadrupoles (3)
+
-
+
-
+
-
+
-
+
+
The ion oscillates in the volume between the quadrupoles
A negative potential at the exit end of the quadrupole can the be used to pull the ion through the quadrupole

## Slide 17
When RF and DC are applied together, Quadrupoles function as a mass analyzer.
The RF-DC voltages can be scanned
  DC voltage increases with time
  RF amplitude increases with time, but frequency remains constant
For one RF-DC combination, only an ion of one particular mass will have a stable trajectory through the quadrupoles.
Technically, it is the mass/charge (m/z) that is affected
Low mass
High mass
Time
Voltage
DC
RF
Mass Spectrometry
Mass analysis by the quadrupoles (1)

## Slide 18
RF-DC combinations in the quadrupole can be used to filter ions by mass:
  ⤷ Low RF-DC combination, ions of low mass transmitted.
  ⤷ High RF-DC combination, ions of high mass is transmitted
Mass Spectrometry
Mass analysis by the quadrupoles (2)
Time
Voltage
m1
m2
m3
m4
m5

## Slide 19
Quadrupoles can function in two modes
  ⤷ Fixed RF-DC : Only ions with a fixed mass will be transmitted
  ⤷ Switching DC off, i.e. RF only:
    ⤷ All ions oscillate within the quadrupole
    ⤷ By keeping a potential at the other end of the quadrupole, all ions can be transmitted together through the quadrupole
Mass Spectrometry
Mass analysis by the quadrupoles (3)

## Slide 20
Mass Spectrometry
Detection of ions
Electron Multiplier
⤷ The sides of the detector are coated with material that produces electrons on impact by charged particles
Detectable Current
+

## Slide 21
Multiple Reaction Monitoring
Based on the theory that, under the same conditions,
  An ion from an analyte will always fragment in the same way
  Fragmentation pattern wIll be unique
Highly selective and sensitive
  We can choose to look at only molecules of interest
  Quantification in ppb levels
Mass Spectrometry
Putting everything together

## Slide 22
Mass Spectrometry
Putting everything together (2)
Three Quadrupoles:
  First Quadrupole will select the parent ion MH+
  Second Quadrupole functions as the collision cell:
    N2 is infused into the collision cell
    Daughter ions produced by Collision Induced Dissociation (CID)
  Third Quadrupole scans the daughter ions
Linear ion trap:
  Second quadrupole ejects all ions except the most intense ion
  Can study how this ion behaves with time
  Trapped ion can be fragmented further by CID
We have AB Sciex Qtrap® instrument at QEL Kochi but we are not using it in routine analysis
N2

## Slide 23
Mass Spectrometry
Putting everything together (3)
Ionization
IQL1
Collision Cell
Q2
IQL2
IQL3
Q1
Q3
Detector
M
+
Select the parent ion
Of interest: Fixed RF:DC
Collision Induced Dissociation
RF Only: Transmit
+
+
+
RF/DC Fixed – Only
One daughter ion
+
+
+
+
+

## Slide 24
Mass Spectrometry
§ in practice

## Slide 25
Multiple Reaction Monitoring
Quantitative analysis using MS/MS
LC MS/MS
M1 + H+
[M1H]+
N2
D1
D2
D3
.
.
Count
“Transitions”
[M1H]+  D3
Multiple such reactions, or transitions, can be anaysed in a single run – hence the name ‘MRM’

## Slide 26
Multiple Reaction Monitoring
Quantitative analysis using MS/MS (2)
In MRM, we choose 2 transitions per analyte
  ⤷ Higher intensity transition is called the quantifier, and is used for quantification
  ⤷ Lower intensity transition is called the qualifier
  ⤷ Quantifier : Qualifier ratio should match between standard and sample
 | Compound | Parent Ion | Daughter Ion | Type | Ratio
LC MS/MS | Ethion | 385 | 199 | Quantifier | 28.6
 |  | 385 | 171 | Qualifier | 

## Slide 27
MS/MS scan modes
Not just MRMs
Q1
Q2
Q3
Detector
M1
M2
M3
M1
d1,d2,d3
Product ion scan
To understand the ways in which an ion can fragment. First step in MS method development
d1,d2,d3
Fixed m/z
Scan
M1
M2
M3
M1
d1,d2,d3
d1
Selected ion monitoring
MRM mode,
Mainly used for quantitation work
Fixed m/z
Fixed m/z
d1,d2,d3
Precursor ion
Scan
To find which ions (precursors) produce a given daughter ion
M1
M2
M3
M1,M2,M3
d4,d5,d5
d1
Scan
Fixed m/z
d1,d2,d3
M1
M2
M3
Scan
(only ions that lost an x fragment)
M1,M2,M3
d1-x
d2--x…
Neutral loss
Scan
To study reaction mechanisms
d4,d5,d5
Scan

## Slide 28
Topics that remain to be covered
in LC-MS/MS
Mass accuracy and resolution, calibration of mass-response, MS tuning
  ⤷ Will be covered in the method validation class
Basic troubleshooting in LC-MS/MS
  ⤷ There will be a separate class on this

## Slide 29
More to learn..
Next session 10th June 2023
(1) Introduction to instrumentation: LC-MS/MS, 03-Jul-23
(2) Introduction to instrumentation: GC-MS/MS, 10-Jul-23
(3) Pesticide residue analysis - Introduction, 17-Jul-23
(4) Advanced pesticide residue analysis, 24-Jul-23
(5) Method validation: requirements and practice, 31-Jul-23
(6) Introduction to measurement uncertainty calculation, 7-Aug-23
✅
🕑
🕑
🕑
🕑
🕑

## Slide 30
Thank you!
§ Questions?
Scientist C, Spices Board
Dr Ramesh BN
WhatsApp
Mobile
Email (O)
Email (P)
Course website
:
:
:
:
:
+91 8547662471
+91 8547662471
rameshbabu.n@nic.in
rameshbn1@gmail.com
http://ramesh.-sb.github.io